DUAL POLARIZATION SPECTROGRAPH

[image: image40.emf]A-B in dB vs input temperature (measured at 24.5 MHz)

-0.60

-0.40

-0.20

0.00

0.20

0.40

0.60

0.80

1.00

1.20

1.00E+04 1.00E+05 1.00E+06 1.00E+07 1.00E+08

Input Temperature at antenna terminal

A-B in dB given average sensitivity = 49 mv/dB


Hardware Manual

RF Associates LLC

February 2013

The Dual Polarization Spectrograph (DPS) is designed to operate between 17 and 34 MHz.  The instrument comprises two identical receivers sharing a common computer controlled, step frequency, local oscillator. The local oscillator operates in the range of 27.7 to 44.7 MHz.  The image frequency range is from 38.4 to 55.4 MHz.

Each receiver employs a 10.7 MHz center frequency, 30 kHz wide IF amplifier providing approximately 49 dB of gain. 
Following each IF strip is a sensitive, wide dynamic range log-detector that develops an output voltage proportional to the log of the input power.  The detector develops a DC output voltage that changes by 25 millivolts for each 1dB change of input power.  A low pass filter and adjustable gain DC amplifier follow the detector.  The resulting signal from each receiver detector feeds a separate 12-bit analog to digital converters (ADC).
The ADCs are read synchronously by a PIC that also generates commands for the direct digital synthesis (DDS) local oscillator, and pulses for the blanking gate circuits which inhibit signals to the ADCs during the time when the DDS is changing frequency.

The dual spectrograph could be fed with RCP and LCP antennas or it could be fed with orthogonal linear antennas.  In the later case a polarization hybrid operating at 10.7 MHz could be inserted between the IF strips and the detectors.  Either configuration could produce RCP and LCP signals at the detector inputs.  

Documentation will simply label the two LO sharing receivers as (A) and (B).  Each receiver uses identical modules (M1 - M4) as seen in Figure 1.  The local oscillator signal generated in M5 is routed thru filters and a power splitter (M6 - M8).  The module interconnect diagram is seen in Figure 2.
The DPS requires a 15 volt DC power supply (15-18vdc) at 500 ma.  
[image: image1.png]



Figure 1.  DPS block diagram
M1  34 MHz Low Pass Filter
This 9th order 34 MHz low-pass filter is designed to minimize the effect of image response signals in the 38.4 to 55.4 MHz frequency range.
M2  RF Front End – HPF, RF Amp, Mixer and Diplexer

The front-end module contains a 17 MHz high pass filter, a 10.7 MHz notch filter, an MAR-3 RF amplifier, a mixer, and a diplexer to provide a broadband impedance match to the mixer output.
M3  10.7 MHz IF Strip

The single conversion spectrograph receiver uses a 10.7 MHz intermediate frequency (IF) amplifier. The IF amp uses 8 poles of monolithic crystal filtering to shape the 30 kHz wide response.  Gain is approximately 49 dB.
M4  Log Detector and Blanking Gate

The detector uses an Analog Devices AD8307 Log Detector followed by a low-pass filter and an adjustable gain DC amplifier.  A 2N7000 MosFet blanking gate inhibits signals from reaching the ADC during the interval when the DDS is changing frequency.  The blanking gate is controlled by a pulse train from the PIC microcontroller.
M5  PIC/DDS/ADCs

This module contains a16F873 microcontroller that communicates with the main control computer via an RS232 link. [Note – do not use a 16F873A in place of the 16F873 as the 16F873 code and the Microchip PICkit3 programmer will not work correctly.]  The PIC generates commands to set the frequency of the DDS-60 DDS oscillator, generates gate pulses, and synchronously samples a pair of MAX 187, 12-bit ADCs. The ADCs digitize the analog signal outputs of the A and B receiver detectors.  The DDS signal is routed thru low pass and high pass filters before being power split to feed the two receiver mixers.
M6  26 MHz HPF

The 26 MHz  HPF is used to minimize signals below 26 MHz generated spuriously in the DDS oscillator.
M7 46 MHz LPF

The 46 MHz LPF is used to minimize signals above 46 MHz generated spuriously in the DDS oscillator.
M8  Power Splitter

The Mini Circuits 2-way zero degree power splitter is used to route the local oscillator signal to the receiver mixers (in M2A and M2B).

[image: image34.jpg]i sams s 15
JENDC » GND

PIC/ODS

DPS DC Power
DisTribuTion




[image: image35.png]pre— Points: 200 IFBiY: 500 Hz AVG: 272 Power. Low
Bias Tee O Reference Plane P1: 0 nm
TRT: 521 D50
Log Mag
Smooth 0 %
CAL: ON (0K)
5.00 dB/ 200
Ref 0.00 4B
150
100
50
- -
00
50
100 \
150 / \
200 / \
0583 536 MHz TR 70775 757 MHz
MET TRALL(TT): 10,699 399 MHz LM 365 08 MK2 TR1: 10.715 MHz LMB.68 0B

MK3 TRT

10.687 MHz,LM:6.50 0B



[image: image36.png]


[image: image37.jpg]S Mz LPr

LH
%D»n\ri 500n\\ 516 n\\ 5a0nY ?.BOY\H g

ct o d C3 C‘l C,b‘ce C7 CB
125 of !5\?‘5 \5\3:\“ ’3517*

cl 07 =\Ro¢F

G2 6= 120 +\8 ¢

g CL 06,0 =327k

1B 28O, 7turns #2/ 00 T50-6 (6)
L2 ,L4= Seoni, jaturns # 2] 80T 5 -6 (8"
L3 = 516, j0 tuvns # 2/ onTS0-6 (8°)

J1.32 BNC



[image: image38.emf]Frequency Response A and B Channels

2280

2300

2320

2340

2360

2380

2400

2420

2440

2460

18 19 20 21 22 23 24 25 26 27 28 29 30 31 32 33

Frequency MHz

ADC Input millivolts

A Channel ADC input Vs Freq

B Channel ADC input vs Freq

1 dB (49 mv/dB)

[image: image39.emf]Channel A and B response at 24.5 MHz to antenna input temperature

1500

1750

2000

2250

2500

2750

3000

3250

3500

3750

4000

1.00E+04 1.00E+05 1.00E+06 1.00E+07 1.00E+08

Antenna input Temp

ADC inputs in millivolts

A Channel

B Channel

























Figure 2.  RF and Control Interconnect diagram
Green = signal, side A
Blue = local oscillator,
Red = signal, side B

Black = gate pulses
Not shown are 3dB pads located at the LO input ports of the A and B channel RF Amp/Mixers.
Power Distribution

The DPS nominal input voltage is 15 vdc (connected to TB1 – right center of drawing).  The allowable range of input voltages is 14.5 to 18 vdc.  At 15 volts the DPS draws approximately 500 ma.  TheDPS is not fused, however steering diodes in the power filter boxes and M5 provide polarity protection. The DPS comprises three 19” rack panels.  The top panel is the “A channel” receiver, the middle panel contains the PIC / DDS/ ADC module and the local oscillator chain filters.  The bottom panel is the “B channel” receiver.  The A and B receivers each have a power filter box which includes a 12 volt three-terminal regulator providing 12 volts to the M2, M3, and M4 modules.  Input power is attached to a barrier strip on the middle panel.  Power is then distributed to the A and B power filter boxes and the PIC/DDS/ADC module (M5).  Module M5 also includes an internal 12 volt three-terminal regulator.  Twelve volt lines from each power filter box are attached to barrier strips for distribution (TB2 for the A receiver and TB3 for B).  When 15 volts is applied to the input barrier strip the local oscillator receives power. Normal current to M5 is 140 ma. Each power filter box has a toggle switch to control power to its receiver.  The ability to turn each receiver on and off may be useful for testing purposes.
[image: image2.png]



           DC Power filter board – 12 volt regulator, steering diode, and RF chokes
[image: image3.png]sPec PR § © §
10,2011 [





            DC Power filter board – trace side
[image: image4.jpg]RFC

Rre ¥ %

L)
N

T o

rpe F O

\2vpC
T




M1A, M1B   34 MHz LPF
[image: image5.jpg]





[image: image6.jpg]Poins: 1772

Bias Tee Off

Log_Mag
Smooth: 10%
58/
Ref0 B

&-TR2 511
SWR
Smooth: 10%
02/

Ref 1

1FBY:TKHz

MI1A 34 MHz LPFF

AVG: 272

PonerHigh

RT\EHE P1: 0mm P20mm

AN

2

T

)

%

<)

=

i)

®

Ik

FiHz_

Az

77

e

T

2z

z

ik

i

i

i

‘5997996 Mz

T

TWHT





M1A 34 MHz LPF response (Top panel insertion loss, bottom panel SWR)
[image: image7.jpg]Poins: 1772
Bias Tee Off

Log_Mag
Smooth: 10%
58/
Ref0 B

&-TRZ 611
SwR
Smooth: 10%
02/

Ref 1

1FBY:TKHz

.

MIB 34 MHz LPF Response

AVG: 272

PonerHigh
e Plane P1: 0 mm P20 mm

T

T

)

=

(e

®

)

®

9979 Mz

I e

I





M1B 34 MHz LPF response (Marker frequencies are shown in boxes just above
the vertical marker lines.)
M2A, M2B  17 MHz HPF, 10.7 MHz Notch, RF Amplifier,  Mixer, Diplexer
[image: image8.jpg]



[image: image9.emf]
[image: image10.jpg]Poins: 1772
Bias Tee Off

Log_Mag
Smooth: 10%
58/
Ref0 B

WA
Smooth: 10%
02/
Ref 1

Channel antenna input to RF amplifier input
Pz

AVG: 272

PonerHigh
vence Plane P1: 0 mm P20 mm

-

.

75

e

7z

z

ik

ik

ik

i3

EEEIEELLS





(A) channel antenna input (M1) to MAR-3 RF amplifier input – C7, with 

MAR3 removed.
[image: image11.jpg]Poins: 1772
Bias Tee Off

Smooth: 10%
58/
Ref0 B

B-TR2 511
SWR
Smooth: 10%
02/

Ret 1

B Channel antenna input to RF amp input

FEW Tk A 22 Powerbigh
R erce Plare P10 P20

i

&3

EEEEEILR THZ TRz





(B) channel antenna input (M1) to MAR-3 RF amplifier input – C7, with 

MAR3 removed.

[image: image12.jpg]1FBY:TKHz

AVG: 272

PonerLow
Reference Plane P1: 0 mm P20 mm

oits: 400
ias Tee Off
= T

Log_Mag

smooh 0% |5

58/

Ref0 B 1

o TR 511 =

SWR

smooh 0% |20

02/

Ref 1 =
<)
®
W

ELL

75

75

=

7z

ik

i3

u/

5Tz

T

Z

ELL





            M2A diplexer response (C10 to IF output – mixer removed)
[image: image13.jpg]M2B Diplerer

Poirts: 400 1FBY:TKHz AVG: 272 PonerLow

Bias Tee 01 695 1z Reference Plane P1: 0 o P20
TR

Log_Mag
Smooth: 0%
58/
Ref0 B

8-TRZ 511
SWR
Smooth: 0%
02/

Ret 1





            M2B diplexer response (C10 to IF output – mixer removed)
M3A, M3B  10.7 MHz IF Amplifier (30 kHz BW)

[image: image14.jpg]



[image: image15.emf]
     10.7 MHz, 30 kHz bandwidth IF amplifier (Gain = 50 dB nominal)
[image: image16.emf]
     Crystal filter impedance matching network

[image: image17.png][e—Tr IFBiY: 500 Hz Power. Low
Bias Tee O Reference Plane P1: 0
TRT: 521 D50
Log Mag
Smooth 0 %
CAL: ON (0K)
5.00 dB/ 200
Ref 0.00 4B
150
100 MK
50 /
- -
00 /
50
100 ]
150 / \
200 / \
0583 536 MHz TR 70775 757 MH

MKT TRALLTT): 10.699 993 MHZ LM7.45 68 MK2 TRT: 10.715 MHz,LM5.18 0B

MK3 TR1: 10,667 MHz,LM:5.30 dB






[image: image18.emf]IF Amplifier Gain Response

-45

-40

-35

-30

-25

-20

-15

-10

-5

0

5

10

15

-90 -85 -80 -75 -70 -65 -60 -55 -50 -45 -40 -35 -30

10.7 MHz Input in dBm

10.7 MHz Output in dBm

B Channel

A Channel


Amplifier (A) gain =49.5 dB, Amplifier (B) gain = 48.7dB.  Compression begins at             about -40 dBm input. 
M4A, M4B Log-Detector, DC Amplifier, Gate



[image: image19.jpg]2N 7660





Overall detector schematic
[image: image20.jpg]Gafl
Qo ra \

Fram
Pin A
ADE3OY
Gam= 1 + K=
R +Rs

e R#SO) =39
Ry=25K,G= /.8

0«’/7% @&é U/.S(? o/lsro:;wecfecf ondl
AD8307//? c/7L fermindled 14 5o
&djusT Rﬁ/ Lar Hooms avlpel
Cuse (Figel scope o7 FluKe 83)

003

LT SN

psec

/

3 Bps k-
R B

Sp €<‘_77051 V’;{gpﬁ Dele

lor




Detail of LM358 Op amp and gate circuit
M5 PIC / DDS / Dual ADC

[image: image21.jpg]el ———T S

@ 7an





M5  PIC / DDS / Dual ADC
[image: image22.png]Output HPF (optional)
P — € DDS 1 otpto (lmo)[)o
¢ Meo 1BRkoogeo0l E] E] E] 9
S eoGnd T | CHENED: s &
o 1BEIogol 4 =
=y Reset. | o Gate — ) XM
il $380808l0og38800 pB8b00808
O 0 Okt C e——— MAX-232
||| eo lDOOOOO00000000/‘1!}000000000
- It °7-|o -3 v
5V Reg l = tal
O | pullasy JE RS-P3p
% < B R S PIC » DDS # ADC -
Lol O O O 00O
o ) ol s o] 18/11/120 5
S o O o
GQ OALYGO OA2 RF
- .ADC = .ADC -





[image: image23.jpg]TOP VIEW

L8] Voo
[15] eno

i
A

MAX2324 [ o
[11] Tiw
[10] T2
8 3] Roour

DIP/SO

EIEF FIE FIR ]

CAPACITANCE ()
DEVICE G 2 03 04 G5
MAX220 0,047 033 033 033 033
MAX232 10 10 10 10 10
MAX232A 01 01 01 01 0f

Diagrams continued in the full data sheet.

55V
VOLTAG!

TIUEMOS
INPUTS

L/CMOS
UTPUTS

To PIC
pin18

pin3





[image: image24.jpg]Poirts: 1000 AVG: 272 FowerHigh

i Tee 1 s | TR Ficerce et 0 P20mn
Log Mag /
o107 [/
9
P
SRS
SwhR 5
S 10
02/ =
it
e N
=
®
©
®
TTMHz TRT B MHz
5
&
=
=
Zz \
i |
s
7
IE] \
‘ —
iz /

OWAz TRZ W




26 MHz HPF Response – this is the HPF in M5 at the output of the DDS.

M6, M7, M8 LO Chain Filtering and Power Splitter
[image: image25.jpg]



M6  26 MHz HPF
[image: image26.png]J1 ¢ c3y cs 7 bl
2EN i \J’ %
= A5 aded A5ad
HPE u ey Laﬂ
I L2 L%
JL 25700 Bl 257k Fe

%I—T LR ! f—ﬁJﬁﬁLﬁo—'«i
T3 ?; iy

27 -7 0 327 106 327
oot c3 ci cs ¢k
He mHz LPF





[image: image27.jpg]ok 1000, IFBW:1kHz, AVE: 22 PowerHigh

Bias Tee Dif ps | i | is | Reference Plane P1: 0 mm P20 mm
o 1

Log_Mag
Smooth: 10% |5
58/
Ref0 B i
&-TR2 511
WA i
Smooth: 10%
02/ 3
Ref 1 —
<)
=
W
®
TOHHz TRT - 0z

7F

75

i |
- |

ik

i3

ik

i e —

ToVHE THZ E0GH





            26 MHz HPF Response
[image: image28.jpg]



M7  46 MHz LPF
[image: image29.png]J1 ¢ c3y cs 7 bl
2EN i \J’ %
= A5 aded A5ad
HPE u ey Laﬂ
I L2 L%
JL 25700 Bl 257k Fe

%I—T LR ! f—ﬁJﬁﬁLﬁo—'«i
T3 ?; iy

27 -7 0 327 106 327
oot c3 ci cs ¢k
He mHz LPF





[image: image30.jpg]Points: 1000 IFBW:1kHz AVE: 22 PowerHigh
Bias Tee Off [THz bs | b3 | s | Reference Plane P1: 0 mm P20 mm
T

Log_Mag

Smoot 10% |5

58/

Ren e |70
B TR2 811

Swh 8

Sk 10%

02/ 20

Ref 1





            46 MHz LPF response
[image: image31.jpg]:ﬂ‘"‘i ‘”D”;' 6 MHz HPF + 46 MHz LPF

Log_Hag
Smooth: 10%
58/
Ref0 B

B-TRZ 511
SWR
Smooth: 10%
02/

Ref 1

IFBW:TkHz

_m

AVG: 272

PowerHigh
Reference Plane P1: 0 mm P20 mm

P

2

T

)

%

<)

®

WAz





            Response of 26 MHz HPF and 46 MHz LPF combined

[image: image32.jpg]Paints: 1000
Bias Tee Off

Log_Hag
Smooth: 10%
58/
Ref0 B

&-TR2 511
SWR
Smaoth: 10%
02/

Ref 1

IFBW:TkHz

AVG: 272

PowerHigh
Reference Plane P1: 0 mm P20 mm

m

2

T

)

%

<)

=

)

®

TOHHz

75

P

=

2z

T

7]

ELLH





            Response from output of DDS to input of power splitter – this includes 

            the 26 MHz HPF on the PIC/DDS/ADC board and the external 

            26 MHz HPF and 46 MHz LPF.

Overall Performance Measurements and Parameters
RF stage (M2)  Gain = 7 dB from RF in to IF out.  This is consistent with the gain of the MAR-3 amplifier specd at 13.5 dB (at 50 MHz), and the TUF-3 mixer with a loss of 5 dB.  MAR-3 noise figure spec is 3.7 dB (at 50 MHz).

IF Amplifier (M3)  Gain = 49 dB.  NF of 2N5109 transistor =3dB. NF of IF approximately 4 dB.  
Spectrogaph NF theoretical is 5 dB (temp = 625 kelvin)

With a galactic background temperature of 25kK at the input of M1 the signal level in a 30 kHz BW is -110 dBm.  At the 10.7 Mhz IF output of M2 this level is 130 kK 
(-103 dBm).  The IF (M3) begins to go into saturation at an input level of -40 dBm which is 63dB above the background noise level at the input of M3.

The galactic background level present at the input of the detector is approximately minus 54 dBm, equivalent to 1E10 kelvin.
With the RF module (M2) IF port terminated in 50 ohms and a 33 MHz LO signal at +7 dBm applied to the mixer, the LO signal level at the RF input port of M2 is -44 dBm.  When the LO is sweeping in the range 27 to 34 MHz this signal level will be present at the antenna terminals of the spectrograph.  The 34 MHz LPF (M1) will attenuate leakage of LO signals above 34 MHz.
DPS Performance Summary

Detector Sensitivity Setting

An HP8640 signal generator (10.700 MHz) followed by a Kay step attenuator was fed into each detector input and the detector gains were adjusted to yield the same DC output voltage to the ADCs. The generator output level was fixed at -40 dBm and the Kay attenuator set to -20 dB yielding an input signal to the detectors of -60 dBm.  Gains were adjusted to yield 1763.8 mv out to each ADC. No splitter was used – the detector inputs were fed directly from the Kay attenuator.  RG58 cables running from the generator to the Kay attenuator and from the attenuator to the detector each had clamp on ferrite chokes.  Next the 10.7 MHz signal level was varied in 5 dB steps as seen in the following table. The A channel detector sensitivity averages 50.65 mv/dB while the B channel detector sensitivity averages 48.32 mv/dB. (Table 1)

	-70
	1251.5
	1276.9
	
	
	
	
	

	-65
	1502.4
	1517.2
	
	50.18
	48.06
	2.12
	

	-60
	1763.8
	1763.8
	
	52.28
	49.32
	2.96
	

	-55
	2015.2
	2002.9
	
	50.28
	47.82
	2.46
	

	-50
	2263
	2241.7
	
	49.56
	47.76
	1.8
	

	-45
	2523.5
	2489.8
	
	52.1
	49.62
	2.48
	

	-40
	2770.9
	2726.6
	
	49.48
	47.36
	2.12
	

	
	
	
	
	50.65
	48.32
	2.32
	Averages

	10.7 MHz
	A volt
	B volt
	
	A chan
	B chan
	A-B 
	

	dbm in
	
	
	
	mv/dB
	mv/dB
	mv/dB
	

	
	
	
	
	
	
	
	


Table 1 Detector sensitivity measuremets

After setting the detector sensitivities to match as close as possible, and in particular to yield the same output voltage at a level expected from the galactic background, the following measurements were obtained. 
Channel A and B, ADC Input Voltage vs Antenna Terminal Temperature

The antenna input of each channel of the spectrograph was connected to the 72 M Kelvin HP461 noise source in series with a Kay step attenuator.  No power splitter was used – direct connection to the Kay attenuator.

The voltmeter option in the DPS software was used to obtain the input voltage to the ADCs as a function of input temperature to each receiver.  This measurement was made at 24.490 MHz. (the frequency read by the voltmeters).  The sweep range of the spectrograph for this measurement was 18-30 MHz.

The attenuator was varied from 40 to 0 dB in 5 dB steps.  At each setting three readings of voltage were obtained and the three numbers averaged to yield a single value for each temperature setting.  Figure 2 shows the close tracking of the two channels as a function of signal input temperature over a 40 dB range.  Channel (A) levels are slightly higher above 30 kK, diverging more from the (B) channel at higher signal amplitudes.



Figure 1. ADC input voltage in millivolts vs antenna input temperature in Kelvin.
In order to get a more detailed view of the divergence between the two channels the difference (expressed in dB) was calculated assuming an average detector sensitivity of 49 mv/dB.   (Figure 2).



Figure 2.  A-B difference in dB vs antenna terminal temperature.  The difference is close to zero at the expected level of the galactic background and reaches a maximum of 1 dB near 10 M Kelvin.

ADC Input Voltage vs Frequency for a constant antenna terminal temperature.

A measurement was made of ADC input voltage vs frequency using a constant input temperature of 228kK.  The voltmeter frequency was varied in 1 MHz steps from 19 to 33 dB while the sweep range of the spectrograph was set at 18-34 MHz.  The result is seen in Figure 3

Figure 3   ADC reading for both Channels A and B as a function of frequency.

Figure 3.   ADC reading for both Channels A and B as a function of frequency.

Ideally this response should be flat – it isn’t. There is a gradual drop off in response amounting to about 2.5 dB over the 19 to 33 MHz range.  Both channels do track quite closely – within about a quarter dB.  The hardware reason for this droop is likely a combination of the RF front end gain (including the 34 MHz LPF response) and the LO power which decreases slightly with frequency.

The effect of this frequency response droop is that the higher frequencies of the spectrograms will be visibly different in color, even if the instrument is driven with a constant temperature, frequency independent, noise source.

A detailed view of the M1 34 MHz LPF responses for both the A and B channels (Figure 4) shows that approximately 1.75 dB of this droop can be explained by the 34 MHz LPF response. 

[image: image33.emf]M1 34 MHz LPF Responses

-2.3

-2.2

-2.1

-2

-1.9

-1.8

-1.7

-1.6

-1.5

-1.4

-1.3

-1.2

-1.1

-1

-0.9

-0.8

-0.7

-0.6

-0.5

-0.4

-0.3

-0.2

-0.1

0

17 18 19 20 21 22 23 24 25 26 27 28 29 30 31 32 33 34 35

Frequency (MHz)

Insertion Loss in dB

Channel A

Channel B


Recall that this filter was required to minimize image response problems when the spectrograph is tuned to low frequencies.  So the droop in response was necessitated by the desire to cover both low frequencies – down to 17 MHz and perhaps lower, while also covering high frequencies.

Solution

 The solution to the droop is likely a simple table of frequency dependent correction coefficients which are applied to the ADC data before the offset and gain controls.  Since both channels track each other quite well, a single correction table applied to both channels should be sufficient.

rf    11 Feb 2013
Circuitry details are provided to Dave Typinsky for use in operation and maintenance of the DPS.  This information should not be distributed to others without the express permission of RF Associates LLC.
M8





M7





M6





M5





M4B





M3B





M2B





M1B





M4A





M3A





M2A





M1A





Antenna (B)





Antenna (A)





Computer





12 Bit ADC





Log-detector


DC amp


Gate





10.7 MHz


30 kHz wide IF Amp





10.7 MHz trap


MAR3 RFAmp


Mixer, Diplexer





34 MHz LPF





17 MHz HPF





2-way


Zero degree power splitter





46 MHz LPF





26 MHz HPF





DDS


27.7- 44.7 MHz





PIC





12 Bit ADC





Log-detector


DC amp


Gate





10.7 MHz


30 kHz wide IF Amp





10.7 MHz trap


MAR3 RFAmp


Mixer, Diplexer





34 MHz LPF





17 MHz HPF







































































�





Power Filter Box





Power Filter Box





Antenna





Antenna





PIC/DDS/ADC





34 MHz LPF





34 MHz LPF





IF Out





RF Amp


 Mixer





RF Amp / Mixer





RF In





RF In





IF Out





LO In





LO In





IF Amplifier





IF Amplifier





Detector





Detector





LO Out





Analog





Gate





46 MHz LPF





26 MHz HPF

















PAGE  
5

