FS-200B Radio Spectrograph System

Instruction Manual

December 2003


[image: image1.png]



RF Associates LLC

1721-I Young St

Honolulu, HI 96826

Table of Contents----------------------------------------------------------------------2

FS-200 Radio Spectrograph System Overview------------------------------------3

Analysis---------------------------------------------------------------------------------5

Module Functions and Layout------------------------------------------------------11

M1
18-30 MHz Bandpass Filter------------------------------------------------13

M2
RF Preamplifier--------------------------------------------------------------15

M3
PIC Microcontroller and DDS 1st Local Oscillator---------------------17

M4
1st LO Bandpass Filter and Amplifier-------------------------------------21

M5
DDS LO Amplifier, Low Pass Filter, Mixer-----------------------------23

M6
Diplexer and 70 MHz IF Amplifier---------------------------------------26

M7
1st IF (70 MHz) Crystal Filter----------------------------------------------29

M8
2nd Local Oscillator (80.7 MHz) and Mixer------------------------------31

M9
Diplexer and 10.7 MHz IF Amplifier-------------------------------------33

M10, M12
10.7 MHz, 4-Pole Crystal Filter-----------------------------------37

M11, M13
10.7 MHz IF Amplifier--------------------------------------------42

M14
Detector-----------------------------------------------------------------------48

Mode Switch/Calibrator Overview------------------------------------------------52

M15
PIC Microcontroller and Driver-------------------------------------------53

M16
Mode Switch-----------------------------------------------------------------58

Power Supply and Distribution-----------------------------------------------------59

FS-200B Radio Spectrograph System

The FS-200B radio spectrograph system includes a radio spectrograph receiver, mode switch/calibrator, computer, controlling software, and power supply.  

Spectrograph

The spectrograph receiver has a pre-detection bandwidth of 30 kHz and steps thru 200 channels in the frequency range 18 to 30 MHz ten time each second.  A wide dynamic range log-detector develops an analog voltage proportional to signal strength as the spectrograph steps thru each channel.  This analog voltage is digitized and passed to the computer for processing and display.  The spectrograph frequency is set by a direct digital synthesis (DDS) oscillator controlled by a microcontroller (PIC), which communicates with the main computer via an RS232 serial link.

Mode Switch/Calibrator 

The mode switch selects the RF source for the spectrograph – for the Florida installation this is either RCP, LCP or a calibrator.  The calibrator comprises a high temperature RF noise source followed by a commercial programmable attenuator.  The step attenuator and mode switch are controlled by a microcontroller (PIC), which communicates with the main computer via an RS232 serial link.

Computer

An HP a300y desktop computer with a 2.4 GHz processor and 512M of RAM is used to control the FS-200B system.  An expansion card has been added to the computer so that it has two serial ports – one for communications with the DDS PIC and the second for communication with the Mode/Cal PIC.  The a300y runs custom spectrograph software that allows remote control of the instrument as well as streaming of spectrograph data   over the Internet.  Connectivity to the observatory LAN is via Ethernet.

Spectrograph Software

Custom spectrograph software running on the a300y computer communicates with both PIC microcontrollers and also generates a local spectrogram display, streams data to the Internet for remote viewing using custom client software, and provides an interface for remote control of the instrument.

Power Supply

An Elenco Quad power supply provides +5, +12, -12 and +15 volts to the spectrograph system through two power cables – one connecting to the spectrograph receiver and one to the mode switch/calibrator assembly.  

A block and signal flow diagram of the FS-200B system is seen in Figure 1.  Functional modules are designated as M1 thru M16 – these are individual modules mounted on two 19” rack panels.  The spectrograph receiver itself comprises modules M1 thru M14 while the mode switch/calibrator includes M15 and M16.



Figure 1   FS-200B Radio Spectrograph functional modules and signal flow

Signal Flow

The FS-200 radio spectrograph is a dual conversion superheterodyne design that operates in the range of 18 to 30 MHz.  High level (+17dBm) mixers provide strong signal handling capability.  

A direct digital synthesis (DDS) oscillator operating between 40 and 52 MHz serves as the 1st local oscillator (LO). This oscillator steps thru 200 frequencies 10 times each second.   The sweep range of the oscillator may be changed remotely via control software so that the spectrograph views different segments of the 18 to 30 MHz range in greater detail.  The number of channels (200) and the sweep rate (200 channels per 0.1 seconds) are fixed.  If a total sweep range of 18 to 28 MHz (10 MHz wide) is selected, then each channel is separated from its neighbors by 50 kHz.  If a sweep range of 6 MHz is selected then the channel separation is 30 kHz. 

The 1st IF frequency is 70 MHz.  A 50 kHz wide, 70 MHz crystal filter protects the second mixer from strong out of band signals.  An 80.7 MHz crystal controlled oscillator serves as the 2nd LO.  The 2nd IF frequency is 10.7 MHz and an IF bandwidth of 30 kHz is determined by a pair of 4-pole crystal filters.

A sensitive, wide dynamic range log-detector develops an output voltage proportional to the log of the input power.  The detector develops a DC output voltage that changes by 25  millivolts for each 1dB of change of input power.  A low pass filter and a DC amplifier with variable gain and offset controls follow the detector.  The resulting signal feeds a 10-bit analog to digital converter.

The log-detector itself  (an Analog Devices AD8307) is extremely sensitive with a minimum detectable signal below –70 dBm.  The 10.7 MHz IF signal delivered to the log-detector from the last IF amplifier is set to approximately –65 dBm.

RF Analysis

The signal flow diagram (Figure 2) shows the distribution of gains and losses within the spectrograph as well as in that portion of the UFRO RF system preceding the spectrograph.  Net gain as measured from the antenna terminals (at the top of a TP antenna) are shown in little “lollipop” circles between each RF stage.  Gains, noise figures, and losses of each circuit element are included in the signal flow diagram.

Signal levels may be calculated at any point in the system given the input signal at the antenna.  For the sake of this analysis it is assumed that the antenna temperature due to the galactic background is 40,000 degrees.  This is equivalent to a –108dBm signal in a 30 kHz bandwidth.  Given this antenna temperature, the output of the UFRO system (the input to M1 - the spectrograph bandpass filter) is –98 dBm with a signal level of –41 dBm developed at the input to the log-detector.   This is almost 25 dB more signal than required at the input of the log-detector and is a result of excess gain being built into the spectrograph.  In operation at UFRO the total spectrograph gain will be reduced either by setting of the step attenuator, insertion of attenuator pads, or removal of the final IF- amplifier stage (M13).  The gain will be reduced to achieve a signal input level of approximately –65 dBm at the input to the log detector.

System Noise Figure 

Noise figure and system temperature calculations have been performed at 3 different reference points using an Excel spreadsheet.

1. Input to the spectrograph bandpass filter M1.  NF = 5.9dB, Tsys=852 deg K

2. Input of the UFRO bandpass filter in the field.  NF = 5.1dB, Tsys = 646 deg K

3. Calculated at TP antenna terminals.  NF = 8.1dB, Tsys=1578 deg K

Receiving system noise compared to the galactic background level is insignificant.


	      Loss 1 in dB
	1
	
	1.259
	
	75.09
	-1
	

	
	
	
	
	
	
	
	
	

	      Gain 1 in dB 
	16
	
	39.811
	
	
	15
	

	         NF 1 in dB
	3.5
	
	2.239
	
	452.2
	
	

	
	
	
	
	
	
	
	
	

	      Loss 2 in dB
	8
	
	6.310
	
	48.69
	7
	

	
	
	
	
	
	
	
	
	

	      Gain 2 in dB
	16
	
	39.811
	
	
	23
	

	        NF 2 in dB
	4
	
	2.512
	
	87.48
	
	

	
	
	
	
	
	
	
	
	

	      Loss 3 in dB
	17
	
	50.119
	
	71.39
	6
	Tsys =

	
	
	
	
	
	
	
	
	852.007

	      Gain 3 in dB
	23
	
	199.526
	
	
	29
	

	        NF 3 in dB
	4
	
	2.512
	
	110.13
	
	

	
	
	
	
	
	
	
	
	NF =

	      Loss 4 in dB
	9
	
	7.943
	
	2.53
	20
	5.95

	
	
	
	
	
	
	
	
	

	      Gain 4 in dB
	23
	
	199.526
	
	
	43
	

	        NF 4 in dB
	4
	
	2.512
	
	4.38
	
	

	
	
	
	
	
	
	
	
	

	      Loss 5 in dB
	3
	
	1.995
	
	0.01
	40
	

	
	
	
	
	
	
	
	
	

	      Gain 5 in dB
	 
	
	1.000
	
	
	40
	

	        NF 5 in dB
	4
	
	2.512
	
	0.04
	
	

	
	
	
	
	
	
	
	
	

	      Loss 6 in dB
	 
	
	1.000
	
	0.00
	40
	

	
	
	
	
	
	
	
	
	

	      Gain 6 in dB
	 
	
	1.000
	
	
	40
	

	        NF 6 in dB
	 
	
	1.000
	
	0.00
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	FS200 System Noise Figure and Temperature as measured at input to bandpass filter M1
	

	
	
	
	
	
	
	
	
	


	      Loss 1 in dB
	0.5
	
	1.122
	
	35.39
	-0.5
	

	
	
	
	
	
	
	
	
	

	      Gain 1 in dB 
	11
	
	12.589
	
	
	10.5
	

	         NF 1 in dB
	3
	
	1.995
	
	323.8
	
	

	
	
	
	
	
	
	
	
	

	      Loss 2 in dB
	6.5
	
	4.467
	
	89.60
	4
	

	
	
	
	
	
	
	
	
	

	      Gain 2 in dB
	16
	
	39.811
	
	
	20
	

	        NF 2 in dB
	3.5
	
	2.239
	
	143.01
	
	

	
	
	
	
	
	
	
	
	

	      Loss 3 in dB
	8
	
	6.310
	
	15.40
	12
	Tsys =

	
	
	
	
	
	
	
	
	646.179

	      Gain 3 in dB
	16
	
	39.811
	
	
	28
	

	        NF 3 in dB
	3.5
	
	2.239
	
	22.67
	
	

	
	
	
	
	
	
	
	
	NF =

	      Loss 4 in dB
	8
	
	6.310
	
	2.44
	20
	5.09

	
	
	
	
	
	
	
	
	

	      Gain 4 in dB
	16
	
	39.811
	
	
	36
	

	        NF 4 in dB
	4
	
	2.512
	
	4.38
	
	

	
	
	
	
	
	
	
	
	

	      Loss 5 in dB
	17
	
	50.119
	
	3.58
	19
	

	
	
	
	
	
	
	
	
	

	      Gain 5 in dB
	23
	
	199.526
	
	
	42
	

	        NF 5 in dB
	4
	
	2.512
	
	5.52
	
	

	
	
	
	
	
	
	
	
	

	      Loss 6 in dB
	9
	
	7.943
	
	0.13
	33
	

	
	
	
	
	
	
	
	
	

	      Gain 6 in dB
	 
	
	1.000
	
	
	33
	

	        NF 6 in dB
	4
	
	2.512
	
	0.22
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	FS200 System Noise Figure and Temperature as calculated at input to bandpass filter in field

	
	
	
	
	
	
	
	
	


	      Loss 1 in dB
	3.5
	
	2.239
	
	359.23
	-3.5
	

	
	
	
	
	
	
	
	
	

	      Gain 1 in dB 
	11
	
	12.589
	
	
	7.5
	

	         NF 1 in dB
	3
	
	1.995
	
	646.2
	
	

	
	
	
	
	
	
	
	
	

	      Loss 2 in dB
	6.5
	
	4.467
	
	178.79
	1
	

	
	
	
	
	
	
	
	
	

	      Gain 2 in dB
	16
	
	39.811
	
	
	17
	

	        NF 2 in dB
	3.5
	
	2.239
	
	285.35
	
	

	
	
	
	
	
	
	
	
	

	      Loss 3 in dB
	8
	
	6.310
	
	30.72
	9
	Tsys =

	
	
	
	
	
	
	
	
	1577.922

	      Gain 3 in dB
	16
	
	39.811
	
	
	25
	

	        NF 3 in dB
	3.5
	
	2.239
	
	45.22
	
	

	
	
	
	
	
	
	
	
	NF =

	      Loss 4 in dB
	8
	
	6.310
	
	4.87
	17
	8.09

	
	
	
	
	
	
	
	
	

	      Gain 4 in dB
	16
	
	39.811
	
	
	33
	

	        NF 4 in dB
	4
	
	2.512
	
	8.75
	
	

	
	
	
	
	
	
	
	
	

	      Loss 5 in dB
	17
	
	50.119
	
	7.14
	16
	

	
	
	
	
	
	
	
	
	

	      Gain 5 in dB
	23
	
	199.526
	
	
	39
	

	        NF 5 in dB
	4
	
	2.512
	
	11.01
	
	

	
	
	
	
	
	
	
	
	

	      Loss 6 in dB
	9
	
	7.943
	
	0.25
	30
	

	
	
	
	
	
	
	
	
	

	      Gain 6 in dB
	 
	
	1.000
	
	
	30
	

	        NF 6 in dB
	4
	
	2.512
	
	0.44
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	FS200 System Noise Figure and Temperature as calculated at antenna terminals
	

	
	
	
	
	
	
	
	
	


Radio Spectrograph Receiver Functional Modules (M1 thru M14)

The spectrograph receiver is a dual conversion design with a 70 MHz 1st  IF followed by a 10.7 MHz 2nd  IF amplifier. The function of each module is described briefly below.

M1 – This 18 to 30 MHz bandpass filter passes the desired band of RF signals from the antenna and provides image rejection and overload protection from out of band signals to the following RF preamplifier stage.  In-band insertion loss is 1 dB or less.

M2 – The broadband RF-preamplifier has 16 dB of gain, a 3.5 dB noise figure and 50 ohm input and output impedance providing a good power match to the input filter (M1) and to the following mixer.

M3 – The PIC microcontroller and DDS oscillator generates the 1st local oscillator signal, which is stepped thru 200 channels in the frequency range 40 to 52 MHz.  The output of the DDS is passed thru a 70 MHz elliptic low pass filter to reduce out of band spurious signals generated by the DDS.  RF output from M3 is at a level of  –9 dBm.

M4 –The DDS local oscillator bandpass filter and amplifier further filters out-of-band oscillator signals and amplifies the desired LO signal to a level of +1.5 dBm.

M5 – amplifier and filter to further condition the 1st LO signal to drive the double balanced mixer (DBM) at a level of +17dBm.  In addition, a 19.2 dB directional coupler is included to provide a sample of the LO signal for test purposes.  With an RF signal from M2 in the range of 18 to 30 MHz and a 1st LO signal in the range of 52 to 40 MHz the 1st IF frequency is 70 MHz.

M6 – A diplexer provides a broadband low VSWR termination for the mixer IF port.  The 70 MHz IF amplifier includes a 6 dB, 50 ohm pad on the output to provide a good impedance match, both to the following filter and also to the IF amplifier itself.

M7 – The 1st IF (70 MHz) crystal filter has an input and output impedance of 50 ohms, a 3 dB bandwidth of 50 kHz, and a center frequency insertion loss of 3 dB.

M8 – The 2nd local oscillator mixer assembly generates a crystal controlled 80.7 MHz signal at a level of +17 dBm.   A 19.2 dB directional coupler is included to provide a sample of the 2nd LO signal for test purposes. With a 1st IF signal frequency of 70 MHz and a 2nd LO of 80.7 MHz the 2nd IF frequency is 10.7 MHz.

M9  - A diplexer provides a broadband low VSWR termination for the mixer IF port.  The 10.7 MHz IF amplifier includes a 6 dB, 50 ohm pad on the output to provide a good impedance match, both to the following filter and also to IF amplifier itself.

M10 – This 4 pole 10.7 MHz crystal filter has a 3 dB bandwidth of 30 kHz. The 5K ohm filter elements are impedance matched to 50 ohms.  Insertion loss is 3 dB.

M11 – A high dynamic range 10.7 MHz IF amplifier followed by a 3 dB pad provides a net gain of 20 dB.

M12 - This 4 pole 10.7 MHz crystal filter has a 3 dB bandwidth of 30 kHz. The 5K ohm filter elements are impedance matched to 50 ohms.  Insertion loss is 3 dB.

M13 - A high dynamic range 10.7 MHz IF amplifier followed by a 3 dB pad provides a net gain of 20 dB.

M14 – The logarithmic detector which has a dynamic range in excess of 35 dB is followed by a 2 kHz low pass filter matched to the 0.5 msec per channel dwell time of the stepped oscillator.  The LPF feeds a DC amplifier with variable offset and gain controls. The DC amplifier output signal, which lies in the range of 0 to +5 vdc, is fed to the DDS PIC (in M3) where it is processed by a 10 bit A/D converter within the PIC.

The spectrograph receiver modules are mounted on a 10.5” by 19”wide rack panel and interconnected using BNC terminated RG-58 coax jumpers.  Wherever possible the modules have been designed with 50 ohm input and output impedances to facilitate testing.  The physical layout of the modules and their interconnecting signal cables is seen in Figure 3.


Figure 3  Spectrograph receiver module locations and interconnecting coax jumpers.

M1

18-30 MHz Bandpass Filter

The 18-30 MHz RF band pass filter includes a low pass and a high pass filter.  Total in- band insertion loss is 1 dB or less, and input and output VSWR is under 1.5:1.


[image: image2.png]


  



[image: image3.png]M:Transmisgion lng Mag 5.8 dB/ Ref B.88 dB &
YA ¥ H 16, @17 MH
dgm A2 4 W E]
=1 VA
-28 <
_25 / 1
~38
-35 //
-43
Abs 7
Start (8.088 MHz St L@, 400 MH=z
b2:Reflectinn SHR g.1 7 Ref %.ﬂﬂ G_,
A Ch2:Hkrh 33847 MH
1.8 f 238
1.7
e f F
o ¢ ]
1.4 ! 1
» ! e N
Abs k-z{ AN
Start L. 0080 Mz 4 Stop 48.808 MHz
1Mk (MH2) dB 21 Mkr (MHZ]
1> 16.82 - I8 e 16.827 §.42
2 18. 4@ -9.88 z: 18.88 1.12
3: 25,84 -8.9¢ 3 25.84 1.27
4: 3d. 40 =1, 12 41 3d.ed 1.89
< 33.65 ~3.97 8> 33.8% 2.98

18 - 30 MHz bandpass filter amplitude response and VSWR




M2

RF Preamplifier

The broadband RF-preamplifier which uses a MiniCircuit Labs AMP-77 device running at 15 volts has 16 dB of gain, a 3.5 dB noise figure and low 50 ohm VSWR providing a good power match to the input filter (M1) and to the following mixer.


[image: image4.png]
















[image: image5.png]O i T

=

p2:Reflecti

nn

,BizTransmisgian 5.0 dB/ Ref 20.98 d8 07
23 fy
15.98 d3

Start 1B.222 MHz Stop 48. 888 MHz

1.@8 c?

N 1 O O

R ]
SN

Start 18.88Q MHz Stop 48. 80¢ MHz

1Mk (MHZ) dB Zi Mk (MHz)

13 75.8h 15.99

RF Preamplifier (AMP-77) Gam and mput VSWR response.




M3

PIC Microcontroller and DDS 1st Local Oscillator

The PIC microcontroller performs several functions.

1. Control the direct digital synthesis (DDS) oscillator frequency

2. Digitize the analog signal strength voltage using its internal 10-bit A/D converter.

3.   Communicate via an RS-232 serial connection with the HP a300y computer

The PIC16F873 microcontroller communicates with the main a300y computer via a MAX232 RS-232 chip. Serial communications rate is 56 K bits, which is derived from an 18.432 MHz crystal.  The 16F873 contains the custom control code for the DDS oscillator and is connected to the DDS board by 4 control lines.  The PIC and MAX232 are mounted on a PC board inside of a small diecast box.  Connections to the DDS board and RS232 9 pin D-connector are via Teflon insulated feedthroughs.  The PIC box is mounted inside a larger diecast box, which also contains the DDS board. 

The DDS board contains an Analog Devices AD9851 DDS and a 70 MHz elliptic low pass filter.   The 180 MHz clock for the DDS chip is derived within the 9851 from a 30 MHz signal generated in an on-board crystal oscillator module.  The DDS board was supplied by Mini-Kits in Australia with the 28 pin 9851 surface-mount DDS oscillator already mounted on the board.  The Mini-Kits board is modified by cutting the trace between pins 22 and 24 of the DDS chip.  A very fine wire is soldered to pin 22 and run to a solder post mounted on the corner of the DDS board.  A control wire runs from the solder post to the PIC enclosure.

CAUTION – when servicing the DDS board do not to break the fine wire running between pin 22 of the DDS chip and the solder post mounted on the corner of the DDS board.


[image: image6.png]




[image: image7.png]



PIC BOX

MAX 232 serial communications and PIC16F873 microcontroller


M4

1st LO Bandpass Filter and Amplifier
The 40 to 52 MHz signal available at the output of the DDS local oscillator module (M3) contains numerous spurious signals lying both above and below the oscillator frequency range.  In order to attenuate these out-of-band DDS oscillator signals module M4 contains a 40-52 MHz bandpass filter. A MiniCircuits ERA-4 MMIC amplifier follows the filter and boosts the LO signal level to +1.5 dBm.


[image: image8.png]ouTPUT

ERA-4
Amplifier






[image: image9.png]P Transmission Log Mag 5.8 dBR/ Ref 15.4d@ dB
_ e Chi:Mkr4 | 70.8dd MH
dB _ [4) [} —41.18 dB
o T 2N
/ AN
-3
g / N
_ /
15 7
=28 7 X
o A ¥ :
Abs ;3 t
Start 18.884 MHz Stop 100.806 MHz
b2:Reflection SHR 8.1 Ref 1.18 ¢z,
i) CHE:HKr1 | 4@. 892 MH]
s ki 1.3
1.9
1.8
1.7
1.6
1.5
1L IAT
1% \
1:3 AW 7iy
Abs 1] 2
Start 1@8.368 MH=z Stop 10@.0600 MHz
1iMkr (MHz) di 2: Hkr (MHz)
i L@. 06 115 1> (7] 1.31
2. 52.8@ 18.77 2: 52 1,32
3: 3d.4@  -25.62 3 3d. 55.085
H> 78,88 -41.18 4: 7d.0d .82





DDS Bandpass Filter and Amplifier Response

M5

DDS LO Amplifier, Low Pass Filter, Mixer

Module M5 contains the first mixer, a MiniCircuits double balanced TUF-1H requiring an LO drive level of +17 dBm. Module M4 delivers an LO signal level of +1.5 dBm so further amplification of the LO signal is required. A MiniCircuits ERA-5 capable of delivering a maximum power level of +18 dBm is used.  Since the ERA-5 has a gain of 19 dB a 3 dB resistive pad is used to adjust the drive level to the ERA-5 so that an output power level of approximately 17 dBm will be achieved.

Following the ERA5 amplifier a MiniCircuits SCLF-65 low pass filter provides additional suppression of spurious high frequency signals from the DDS oscillator.

A MiniCircuits 19.2 dB directional coupler (PDC-20-1W) located just before the LO port of the mixer is included to provide a sample of the LO signal for test purposes.  This test port should be terminated in 50 ohms when unused.  Since the normal LO drive level is approximately +17 dBm the power available at the coupled arm test port is approximately –2 dBm.  

With an RF signal from M2 in the range of 18 to 30 MHz and a 1st LO signal in the range of 52 to 40 MHz the 1st IF frequency is 70 MHz


[image: image10.png]


 DDS LO Amp, Low Pass Filter, Directional Coupler , and Mixer


M6

Diplexer and 70 MHz IF Amplifier

The input diplexer provides a broadband low VSWR termination for the mixer IF port.  

The 70 MHz IF amplifier uses a low noise 2N5109 transistor drawing approximately 50 ma. The 2N5109 is a silicon NPN transistor in a TO39 type package designed specifically for broadband applications requiring good linearity. The noise figure of the 2N5109 is rated at 3 dB at 200 MHz   Transformer T1 is a 4:1 balun. A 6 dB, 50 ohm pad on the output helps to provide a good impedance match to the following filter and also stabilizes the input impedance of the IF amplifier itself. 

Gain from input to output of the module is 9 dB, which includes the 1 dB loss in the diplexer and the 6 dB loss in the pad.  The 2N5109 amplifier itself therefore has a gain of 16dB at 70 MHz.


[image: image11.png]


 



[image: image12.png]dB
1
1

=1
=i
Abs

O e e n e
—~ N o F o

b
@

Bl:Transmission /M Log Mag 5.0 dB/ Ref 25.48 dB C
ChizMkrl 78. #@s MH
g 8.7¢ dB
5 +
a E’\
5
- =
o
S
N
Start 18.0808 MHz Stop 200. 460 7
b2:Reflection SWR Bl / Ref _@.98 a7
:J’J
M
2
r/
;f’ﬂ Vl
M
A"
Start 16.0088 MHz Stop 208. 806 MHz
1iMkr (MHZ) dB 2: Mkr (MHz)
1> 7@.81 B.72





M7

1st IF (70 MHz) Crystal Filter

The 1st IF (70 MHz) crystal filter (Filtronetics FN2009) has an input and output impedance of 50 ohms, a 3 dB bandwidth of 50 kHz, and a center frequency insertion loss of 3 dB.  Ultimate attenuation is specified as 60 dB with a maximum of 1dB of in-band ripple.  The filter bandwidth marked on the unit is 26 kHz but this is actually +/- 26 kHz.


[image: image13.png]




[image: image14.png]>

[ SR

\Pl:Transmission /M Log Mag 5.9 dB/ Ref @.@6 dB

I 3/ | Chi:[Mkr3 | 78.826 MH
=z O =5.72 dB
- /2 \

55
5i / 3

o AN

40] L7 N\

50 | ]

Siart 69,808 Miz Stop 70,206 M2
b2:Reflection SWR a.1 Ref .aa 2
g Ch2:Mkrl | 7. 088 MH

; .28
7
b A
5
y i

= AR
2 T

£

Start 69.800 MHz Stop 70.28@ MHz
[:MKr (MHz) _ dB 21 Hkr (M=)

1T 70.86  -2.82 1> 70.08 T.23

P: £9.97  -5.9@ 22 69973 y.31

3> 7d.43  -5.72 370026 2.54

Filtronetics FN-2009 SN 1847




M8

2nd Local Oscillator (80.7 MHz) and Mixer
An 80.7 MHz crystal oscillator (2N-5179) is followed by an 8 dB pad, a low pass filter (MCL SALF-78), an ERA-5 amplifier, an MCL PDC-20-1W directional coupler, and a MCL TUF-1H high level mixer.  The second LO signal is injected into the mixer at a level of +17 dBm. 

The ERA-5 is capable of delivering a maximum power level of +18 dBm. Since the ERA-5 has a gain of 19 dB, an 8 dB resistive pad is used to adjust the drive level from the oscillator so that an output power level of approximately 17 dBm is  achieved. 

The 19.2 dB directional coupler is included to provide a sample of the 2nd LO signal for test purposes. This coupled port should be terminated in 50 ohms when not in use. With a 1st IF signal frequency of 70 MHz and a 2nd LO of 80.7 MHz the 2nd IF frequency is 10.7 MHz.


[image: image15.png]




M9

Diplexer and 10.7 MHz IF Amplifier

The input diplexer provides a broadband low VSWR termination for the mixer IF port.  

The 10.7 MHz IF amplifier uses a low noise 2N5109 transistor drawing approximately 50 ma. The 2N5109 is a silicon NPN transistor in a TO39 type package designed specifically for broadband applications requiring good linearity. The noise figure of the 2N5109 is rated at 3 dB at 200 MHz   Transformer T1 is a 4:1 balun. A 6 dB, 50 ohm pad on the output helps to provide a good impedance match to the following filter and also stabilizes the input impedance of the IF amplifier itself. 

Gain from input to output of the module is 16 dB, which includes the 1 dB loss in the diplexer and the 6 dB loss in the output pad.  The 2N5109 amplifier itself therefore has a gain of 23 dB at 10.7 MHz.


[image: image16.png]









[image: image17.png]Bbl:Transmission /M Log Mag 5.6 dB/ Ref @.@8@dB C
ChiziMkerl 18. 589 MH
dB T 16.3% o3
15 ]
ia
) e
e e
-5 gy 1
-18 iy
-15
Abs
Start 1.0808 MHz Stop 10@.R288 MHz
B2:Reflectinn SWR a.1 / Ref B8.9@ c
Ch2:Mkrl | 18.589 MH
1.7 1.88
1.6
1.5
1.4
1.3
2R boccor] 2
T A
Abs I
Siart 1,800 MHz Stop 180,288 MHz
1:Mkr (MHz) dB 2i Hkr (MHz)
1> 18.59 16.3% > 18.59 1.83
B:  1¢8.48 -7.%6

10.7 MHz Diplexer, 2N-5109 Amp and 6 dB Pad
Gain and Input VSWR





[image: image18.png]e

2
@

»2:Reflection SHR 4.1 / Ref @.98 62

Ch2:lMkrl 18, 589 MH:
1.86
.7
&5
i P N v
3 PO
2 st
5 _ar]

«l e

Start 1.880 MHz Stop 184.860 MHz

iMKr (HHz) dB 2: Hkr {MHz)

s 18.59 1.6

10.7 MHz Diplexer, 2N-5109 Amp, 6 dB Pad Output VSWR




M10, M12
10.7 MHz 4 pole Crystal Filter

Each IF crystal filter module uses a pair of International Crystal Manufacturing (ICM) 10M-30B matched, two-pole, crystal filters with a 3 dB bandwidth of 30 kHz. Each filter element has a characteristic impedance of 5 K ohms.   Ferrite toroid transformers are used to match the filter impedance down to 50 ohms. The module input and output VSWR is under 2:1.  Typical insertion loss is 3 dB with an ultimate attenuation better than 50 dB. 


[image: image19.png]



Each toroid is mounted to the PC board using nylon screws.  The crystal filters are mounted under the PC board and are not visible in the photograph   Each filter has two ports designated as port (A) and port (B).  For a given filter module, the SWR response is practically identical, regardless of the port being measured.  



[image: image20.png]DB Transmission Log Mag 1 5.8 dB/ Ref @.6@8 dB C?

—~—%— 1 Chi:Mkrl | 1@. 788 MR
dB 7 —2.99 dB
-18
~-15
—on /’ \\
=25 7 X
-38 7 Y
=35 ¥
-4
L {
Abs Ty T
Start 10.5808 MHz Stop 14. HH’}
b2:Reflection SWR a.1 / Ref 1.18 G.?
Ch2:jMkr1 Hﬂ.? A M
- L 1.8}
1.8
e fl \
1.6 T
1.5 i
1.4 Vv
1.3
Abs
Start 10.50@ MHz Stop 18.806 MHz
1:Mkr (MHz) dB 2: Mkr (MHz)
1> 18.78 ~-2.99 1> im.7@ 1.87

FILTER #1 PORT A





[image: image21.png](M Transmission Log Mag 1L 5.6 dB/ Ref @.@8 dB C7
- ChizMkr1 | 18, 788 MH
it V4 2.98 dB
b
o8 // \\
= 7 %
-38 7 <
=33 [
-48 i
Abs = T
Start 1@.808 MHz Stop L@.BB@ M
b2:Reflection SWR a.1 Ref 1.18 07
# ChZ:[Hkr 1. 784 M
1.9 \ 1.9%
e 1
1.7 i
- VAR
: V]
1.4 o
1.3
Abs
Start 1@.400 MHz Stop 1@.860 MHz
1:Mkr (MHz) dB 2: Mkr (MHz)
1> 18.7@ =40 1 18.7@ 1.9%

FILTER #1 PORTB





[image: image22.png]PliTransmission Log Mag L 5.8 dB/ Ref @.48 dB  C7

— [ Chiz[Mkrl | 18, 7@@ MH
dB 7 ~2.90 d8
-18 P
-15
2 7 X
:gg / \\.,J“
-4g : t
ibs / i
STart 10608 Mz Stop 10,808 MR
D2:Reflection SR 2.1 Ref 1.18 C7
¥ @ Ch2:HKkrlf | _1@. 788 M
1.9 2 1.9¥
: A
" (|
% P
L5 YT
. V1
1.4
1.3 v
Abs
Start 18,488 MHz Stop 10.886 MHz
(:Hkr (MHz) o8 2 Hkr_(MHz)
1> 18.786  -2.92 1> 18.78 T.97

FILTER #2 PORT A





[image: image23.png](Bl:Transmissinn Log Mag L5.8 dB/ Ref 6.8@8 di C?

—~V— 1 Ghifkrl | 10,700 M
< / ~2.99 d3
18
-15
oa // \\
= 7 N
x 7 N
Y
48 7 |
ibs / \‘
Siart 10.600 Mz Stop 10888 M
D2:Reflection SWR a.1 Ref 18 |8}
T Chz:krl | |£2. 788 M
1.9 \ .9
1.8 il
Lt It
E (Y}
: 1V
1.4 ¥
1.3
Abs
Start 1@.680 MHz Stop 1@.808 MHz
1: Mkr (MH=z) dB 2: Mkr (MHz)
1> 10.768 -2.98 1> 18.7d 1:9%

FILTER #2 PORT B




M11, M13

10.7 MHz IF Amplifier

The 10.7 MHz IF amplifier uses a low noise 2N-5109 transistor drawing approximately 50 ma. The 2N5109 is a silicon NPN transistor in a TO39 type package designed specifically for broadband applications requiring good linearity. The noise figure of the 2N5109 is rated at 3 dB at 200 MHz   Transformer T1 is a 4:1 balun. A 3 dB, 50 ohm pad on the output helps to provide a good impedance match to the following stage as well as stabilize the input impedance of the IF amplifier itself.

Gain from input to output of the module is 20 dB, which includes the 3 dB loss in the output pad.  The 2N-5109 amplifier itself therefore has a gain of 23dB at 10.7 MHz.


[image: image24.png]



.



[image: image25.png]Pl:Transmission Log Mag 2.6 dB/ Ref 24.@@8 dR C7?
Chiz[Mkrl | 1@, 7jdd MH
demm_s 26, 4+ dB
18 e g
14
i
12
18 1
8
hbs
Start 1.088 MHz Stop 180. 0868 MHz
b2: Reflection 3R 8.1 / _Ref 0.80 €7,
Chz:[Hkrl | 10,00 M
AN ]
1.6 s
1.5 Y
1.4 pOL
1_ 3 \ A[\/J
1.2R— 22
1-1 P V!
1
Abs ]
Start 10088 MHz Stop 1@0. 0808 MHz
L:Mkr (HHz) B Z: HKr_(MHz)
15 16.76  28.37 1> 18.78 1.8

10.7 MHz IF AMP #1 GAIN and INPUT SWR





[image: image26.png]N F

Abs

b2:Reflection SWR a.1 7 Ref '6.60 c?
Ch2:Mkr1 1d. 788 MH
1.8p
N ’_‘_,/V"""
Start 1.##8 MHz Stop 1#f. 808 MHz
1:Mkr (HHz) dB 2 Mkr (MHz)
> 878 -29.8% > i@.7d T.%6

10/7 MHz IF AMP #1 OUTPUT SWR





[image: image27.png]Bi: Transmission Log Mag 2.6 dB/ Ref 24.48 dB C7?

i Chi:Mkrl [ 1@.7@d MH:
o i 19.9b dB
18
16
14
12 ]

18
8
Abs
Start 1.0088 MHz Stop 100. 2008 MHz
p2: Reflection SHR 8.1 / Ref 0.80 c?
ChZ:Mkr1 | 1@. 788 MH
1.85
i e
1.4 o a0
1.3 PPV acasn
1.2 s
i 1\ st
(|
Abs
Start 1.088 MHz Stop 180. 7A@ MHz
1:Mkr (MH2) ol 2:Mkr_(MHz)
> 18.78 19.96 1> 108.78 1.65

10.7 MHz IF AMP #2 GAIN and INPUT VSWR





[image: image28.png]b2:Reflectinn SWR a.1 / Ref @.80 c?

Ch2zlMkr1 1d. 748 MH:
1.6 1.12
1.5
14
1.3 P
1.2 AN ]
1:1 S —
1
Abs ?
Start 1.008 MHz Stop 1@@.000 MHz
1: Mkr (MHZz) dB 2: Mkr (MHz)
1> 18.76  -23.4%1 > 18.7d 1.12

10.7 MHz IF AMP #2 OUTPUT SWR




M14

Detector

The logarithmic detector (AD8307) develops an output voltage proportional to the log of the input power.  It is followed by an active 2 kHz low pass filter matched to the 0.5 msec per channel dwell time of the stepped oscillator.  The LPF feeds a DC amplifier with variable offset and gain controls. The DC amplifier output signal, which lies in the range of 0 to +5 volts, is fed to the DDS PIC (in M3) where it is processed by a 10 bit A/D converter within the PIC.  The low pass filter and DC amplifier both use low offset, low drift, JFET input LF411 op amps.

The 8307 is extremely sensitive, with a minimum detectable signal level of below -70 dBm and a dynamic range of over 90 dB.  With its input terminated in 50 ohms, the 8307 output is approximately one hundred millivolts.  The 8307 DC output voltage is proportional to the log of input power. A 1 dB change of input power produces a 25 millivolt change in output voltage.  The maximum output voltage from the detector is 2.5 volts.

As used in the spectrograph, the output of the log detector is low pass filtered to remove noise, and offset to remove the 8307 quiescent output level of approximately 100 millivolts.  Then the signal is amplified to cover the full (0 to 5 volt) range of the A/D converter.  Amplification increases the scale factor from the original value of 25 millivolts per dB to over 100 millivolts per dB.  Amplification of the 8307 signal increases its sensitivity to weak signals but reduces the dynamic range of the detector to less than 35 dB.  A 5 volt zener diode and trimpot at the output of the DC amplifier limit the maximum voltage produced by the detector module thus protecting the A/D converter from over-voltages that could be produced by very strong signals.

Trimpots are used for gain, offset and to set the maximum signal level that is delivered to the A/D converter.  These trimpots are accessed thru holes in the bottom of the diecast box holding the detector.  The box must be removed from the front mounting plate and adjustments made thru holes in the bottom of the box.  The cover of the box cannot be opened for access to the trimpots because the detector is very sensitive to stray fields.  Any adjustments to the trimpots must be done with the box tightly sealed.

Plus and minus 12 volts DC is supplied to the log detector module.  Internal 3- terminal regulators (78L09, 79LO9) develop +9 and –9 volts for the op amps and a 78L05 produces +5 volts for the AD8307.


[image: image29.png]OUTPUT _/






[image: image30.wmf]Log Detector Response

0

1

2

3

4

5

-75

-70

-65

-60

-55

-50

-45

-40

-35

-30

-25

10.7 MHz RF Input (dBm)

DC Output (Volts)


The Log Detector response curves were obtained as follows.

1. Terminate the input to the log detector in 50 ohms and adjust the OFFSET trimpot for zero output.

2. Using a 10.7 MHz signal generator set the output level trimpot so that the maximum output level from the detector module is 5 volts as the generator output is increased above –25 dBm.

3. Set the gain trimpot for the desired response – in the plot above the top curve (with a slope of 150 mv/dB) is maximum gain and the lower trace (with a slope of 115 mv/dB) is with the trimpot set for minimum gain.  The dynamic range for the maximum gain setting is 65-45=20 dB and the dynamic range for minimum gain is 65-30=35 dB.  For normal spectrograph operation adjust the gain trimpot for minimum gain.

The spectrograph gain in front of the detector module is set (using attenuator pads or a step attenuator) so that the input signal to the detector module is approximately –65 dBm on the galactic background.  This corresponds to a DC output from the detector of approximately +1 volt.  The total attenuation used will depend upon the amount of RF and IF gain ahead of the detector module.  Attenuation is normally inserted into the spectrograph at the time of installation and should not need to b e changed.  The location and distribution of the attenuation may affect the spectrograph noise figure and should only be installed or changed after proper analysis.

Notice that this set up procedure for the detector does not utilize the full dynamic range of the A/D converter as the signal level out of the detector is approximately 1 volt on the galactic background.  However if the offset and gain controls were set to give zero volts out on the background and an amplifier in the front end were turned off then the offset would be so large that it would drive the detector module output voltage negative – possible damaging the PIC micro-controller and its internal A/D converter.

Mode Switch/Calibrator 

A mode switch (M16) with 3 coaxial relays (K1, K2, and K3) selects the RF source for the spectrograph (for the Florida installation this is either RCP, LCP or a calibrator).  

The calibrator comprises a high temperature RF noise source followed by a commercial programmable attenuator.  










Both the mode switch and the step attenuator are controlled by a microcontroller (PIC), which communicates with the a300y computer via an RS232 serial link.

The mode switch/calibrator components are mounted on a 5.25” by 19”wide rack panel.

M15 – A PIC microcontroller and driver circuit generate command signals to control a commercial step attenuator and the mode switch.  Three control lines operate the LCP, RCP, Cal/Opr Mode switch and seven control lines drive the individual 1, 2, 4, 8, 16, 32, and 64 dB cells of the step attenuator.  The PIC communicates with the a300y computer via an RS232 line.

The physical layout of the modules and their interconnecting signal cables is seen below.























The PIC controller module is connected to the Elenco Power Supply by a 3-wire cable supplying +5, and +12 volts.

Pushbuttons on M15 can be used to test the RCP/LCP polarization select relays as well as the calibrate function.  One push of the CAL button initiates the CAL sequence while pushing the RCP/LCP button toggles the spectrograph input between the RCP and LCP antennas.


[image: image31.png]



PIC Microcontroller and Driver Module (M15)

Numbers on PIC controller box indicate PIC and MAX-232 pin numbers.


[image: image32.png]



The PIC16F873 microcontroller communicates with the main a300y computer via a MAX232 RS-232 chip. Serial communications rate is 56 K bits, which is derived from an 18.432 MHz crystal.  The 16F873 contains the custom control code for the mode switch and the calibrator programmable attenuator.  The PIC and MAX232 are mounted on a PC board inside of a small diecast box.  Connections to the driver board and RS232 9 pin D-connector are via Teflon insulated feedthroughs.  The PIC box is mounted inside a larger diecast box, which also contains the relay driver board.



Driver circuits.

Two different driver circuits are used – one to control the 5 volt OMRON RF relays in the mode switch (M16) and the other, a 12 volt circuit used to drive the solenoids in the Kay programmable step attenuator.  Each relay solenoid requires a dedicated driver circuit.  The drivers turn on their solenoid when a +5v signal is applied by the PIC.  



M16

Mode Switch

The mode switch contains three, 5 volt, SPDT Omron RF relays.  These relays allow selection of either RCP or LCP signals and also control selection between calibrate and operate modes.  The relays are controlled by signals generated by the PIC and amplified in the driver circuits located in module M15.


[image: image33.png]



Spectrograph System Power Supply and Power Distribution

Power for the FS-200 spectrograph receiver and mode switch/calibrator assembly is provided by an Elenco Quad power supply.  This unit supplies all necessary voltages for the system (+15, +12, -12, and +5 vdc).  The +12, -12, and +5 supply voltages are fixed by internal regulators, while the +15 volt supply output voltage is set using a potentiometer mounted inside the Elenco enclosure.


[image: image34.png]




The Elenco supply has been modified by adding a 5-pin female circular power connector (to provide +15, +12, -12, and +5 volts to the spectrograph receiver), and a 3-pin female power connector (to provide +12, and +5volts to the mode switch/calibrator assembly).  








Two power cables are provided, one with 5 conductors for the spectrograph, the other with 3 conductors for the mode switch/calibrator assembly.

















Low Pass 





High Pass 





30 





30 





100 





100 





5-60





5-60





5-60





5-60





330 





330 





330 





330 





330 





330 





5-60





5-60





5-60





5-60





100





100





220





M 14





M 13





M 12





M 11





M 10





M 9





M8





M7





M6





M5





M4





M3





M2





M1





DDS RF Out





Analog


In








Analog


Out





RF


In








IF


Out





LO


In





RF


In





40 - 52 MHz


BPF and


ERA-4 Amp








RF In





PIC and DDS








DC Power


Filter








Log Det





10.7 MHz  


IF Amp





10.7 MHz  


IF Amp





10.7 MHz


XTAL


FIlter





10.7 MHz


XTAL


FIlter





10.7 MHz 


Diplexer 


and 


IF Amp





80.7 MHz


XTAL Osc


and Mixer





70 MHz


XTAL


FIlter





70 MHz 


Diplexer and 


IF  Amp





DDS Amp


and Mixer





AMP-77


RF Amp





18-30 MHz


Bandpass


Filter





(10.7 MHz IF)





(Analog Signal Strength to A/D)





CAL





LCP





RCP





M16





M15





(RS-232)





Noise Gen





Step


Atten





Mode


Switch





Drvr





PIC





Computer





(RS-232)





(10.7 MHz IF)





(70 MHz IF)





(1st LO Test)





(2nd LO Test)





(40 – 52 MHz)





(RF IN)





M14





M13





M12





M11





M10





M9





M8





M7





M6





M5





M4





M3





M2





M1





DC AMP





LPF





Log Det





Pad





10.7 MHz


XTAL


BPF





Pad





10.7 MHz


XTAL


BPF





Pad





Dpxr





80.7


MHz


XTAL OSC





LPF





CPLR





70 MHz


XTAL


BPF





LPF





DDS





A/D





PIC





Pad





Dpxr





BPF





Pad





RF Amp


AMP-77





CPLR





LPF





18-30 MHz


BPF





220





+15 vdc





GND





RF Out





RF In





+15 vdc





RF Out





RF In





Bottom View





10 fd





0.1





AMP-77





1N-4001





Ferrite Beads





RF Out


-9 dBM


40 to 52 MHz





Ferrite beads





+5 vdc





RS-232





PIC Reset





PIC Reset





Analog In











red





red





22





+5 vdc





purple





gray





gray





brown





green





yellow





white





gray





MAX232 Pins





PIC Pins





DDS Board





PIC Box





XTAL





DDS





Filter





25





7





8





Analog In





5





4





3





2





25





24





23





14





22





8





1





2  





1 











To DDS Pin 22





To DDS Pin 8





To DDS Pin 25





To DDS Pin 7

















To DDS Pin 22





To DDS Pin 8





To DDS Pin 25





To DDS Pin 7





Analog In








+5 VDC








C9


0.1








C7


18 pf








Reset








+5 VDC








C6


0.1








+





C5


100fd








R1 (1K)








C8


18 pf








Y1








C4








C3








C1








C2








+





+





+





+





DB-9





9





8





7





6





5





4





3





2





1





1





2





3





4





5





6





7





8





9





10





11





12





13





14





28





27





26





25





24





23





22





21





20





19





18





17





16





15





16





9





10





11





12





13





14





15





1





8





7





6





5





4





3





2





Reset





Analog In





IC-1





IC-2





FS-200 Spectrograph PIC 16F783 Microcontroller and MAX232





Note:	C1, C2, C3, C4 are 1 fd


	IC1 = MAX-232,  IC2 = 16F-873,  Y1 = 18.432 MHz


	Ground RESET line to reset PIC





IC-2





IC-1





+5 VDC





22





24





26





1





R1





C2





C3





C1





C4





C6





C5





C9





C8





Y1





C7











Pin 3 DB9





Pin 2 DB9





10





8





1





6





20





28











+15 VDC





70 MHz IF Out





Q1





Mixer





Low Pass


Filter





Dir Cplr





4





3





2





1





8





7





6





4





5





3





2





1





2





4





6





8





7





5





3





1





C4





C5





C3





C2





C1





VR1





R4





R3





R2





R1





Note: C1, C2, C3, C4 are 0.1 fd


        *R4 is 68 ohms ½ watt








Bead





R4*


(68)





C3





-19 dB


Coupled


Arm





DDS LO Amp, Low Pass Filter, Directional Coupler , and Mixer





40 to 52 MHz  LO In


+5 dBm





18 to 30 MHz RF In








40 to 52 MHz LO In +5 dBm











C2





VR1








Q1





3 dB Pad





R1, R3


(294)





R2


(17.4)





78L09





C1





C4





+





C5


(10 fd)





+15


VDC





MCL


ERA- 5





70 MHz IF Out








-2 dBm for


+17 dBm


LO Drive Level 





MCL


TUF-1H




















(-19 dB)





MCL


PDC-20-1W








Dir Cplr





Low


Pass





MCL


SCLF-65





18 to 30 MHz RF In

















In





Out





Gnd





Gnd





ERA-5





1  RF








2  IF








3  GND








4  LO








Bottom View





TUF-1H














1








3








5








7








2








4








6








8








Top View





SCLF-65


1 In


8 Out

















PDC-20-1W


1 In


2 Coupled


3 Out








1








2








3








4








5








6








7








8








Bottom View








L1








C1








C2








R1








L2








C4








C3








R2








C5








R3








R4








C6








R6








R5








e








b








c








T1








Q1








R7








R8








R9








R10








C7








C8








C9








+15 VDC








70 MHz


IF Out








70 MHz


IF In








R4


(390)





R3


(2K)





R5


(27)





R6


(5.1)





e





b





c





C5








C6








R8


(150)





Q1


2N-5109





6 dB Pad





R10


(150)





R9


(37.4)





T1*





C7








C8








C9








R7


(27)





+15 VDC





RF Out





RF 


In





R1


(51)








R2


(51)








L2*


(28 nh)








L1


(330 nh)








C1


(4.5-60 pf)








C2


(100)








C4


(2-10)








C3


(100)








Diplexer





70 MHz Diplexer, 2N-5109 Amp, and 6 dB Pad





Note: 	C5, C6, C7, C8, C9 are 0.1 fd


	Q1 requires a heat sink


	*T1 is 10 bifilar turns of #26 wire on Amidon T50-10 core


	*L2 is 3T #21 wire close wound on 1/8” drill bit





R4


(390)





R3


(2K)





R5


(27)





R6


(5.1)





e





b





c





C5








C6








R8


(150)





Q1


2N-5109





6 dB Pad





R10


(150)





R9


(37.4)





T1*





C7








C8








C9








R7


(27)





+15 VDC





RF Out





RF 


In





R1


(51)








R2


(51)








L2*


(28 nh)








L1


(330 nh)








C1


(4.5-60 pf)








C2


(100)








C4


(2-10)








C3


(100)








Diplexer





70 MHz Diplexer, 2N-5109 Amp, and 6 dB Pad





Note: 	C5, C6, C7, C8, C9 are 0.1 fd


	Q1 requires a heat sink


	*T1 is 10 bifilar turns of #26 wire on Amidon T50-10 core


	*L2 is 3T #21 wire close wound on 1/8” drill bit





PDC-20-1W


1 In


2 Coupled


3 Out








1








2








3








4








5








6








7








8








Bottom View








1








3








5








7








2








4








6








8








Top View





SALF-78


1 In


8 Out

















1  RF








2  IF








3  GND








4  LO








Bottom View





TUF-1H














Note: C8, C10, C11, C12, C13 are 0.1 fd





ERA-5





In





Out





Gnd





Gnd





e





b





c





Case (Gnd)





2N5179





Bottom View





+15


VDC





c





MCL


SALF-78





MCL


ERA- 5





Low


Pass





Dir Cplr





MCL


PDC-20-1W








(-19 dB)











C1


(560)





R1


(3.3K)





R2


(2K)





R3


(39)





R4


(510)





C2


(560) 





L1


(69 nh)





R5


(100)





C8





C9


(10 fd)





+





R9


(10)





C10





C5


(5-40)





C3


(18)





C6


(18)





C4


(100)





C7


(100)





C12





78L09





C13





C11





R7


(52.3)





R6, R8


(115)





8 dB Pad





R10*


(68)





Y1


80.7


 MHz





Q1


2N5179





b





e





MCL


TUF-1H








Q2





-2 dBm for


+17 dBm


LO Drive Level 





70 MHz RF In








10.7 MHz IF Out








VR1








R1








R2








R3








R4








R5








R6








R7








R8








R9








R10








Q2








Y1








C4








C7








C3








C6








C5








C2








C1








L1








C8








Low


Pass


Filter











Dir Cplr








Mixer








-19 dB


Coupled Arm








C12








C13








C10








C11








10.7 MHz IF Out








70 MHz RF In








+15


VDC





1








2








3








4








5








6








7








8








1








2








3








4








VR1








80.7 MHz XTAL Osc, Low Pass Filter, Directional Coupler, and Mixer











C2








R2








R1








R3








L1








R5








R4








b








c








e








C1








C3








R9








R8








R7








C5








C4








T1








R6








C6








Q1








RF


 In








RF


Out








+15 VDC








Note:  C2, C3, C4, C5, C6 are 0.1 fd





T1 is 10 bifilar turns of #28 wire on Amidon FT37-43 core (4:1 Transformer)





Heat Sink required on  Q1 (2N5109), current 50 ma














10.7 MHz Diplexer, 2N5109 Amp, 6 dB Pad





R1


(51)





R3


(390)





R2


(2K)





R4


(27)





R5


(5.1)





e





b





c





L1


(330 nh)





C1


(100 pf)





C2








C3








R7


(150)





Q1


2N-5109





6 dB Pad





R9


(150)





R8


(37.4)





T1





C4








C5








C6








R6


(27)





+15 VDC





RF Out





RF 


In





Diplexer





2.1vdc





1.4vdc





13.6vdc





100





100





560 nh





T1





T2





560 nh





F1





F2





T1, T2:  Primary 5T #24, Sec 40T #28, Core FT50-43


F1, F2:  ICM 10M-30B, 10.7 MHz Crystal Filter, 30 kHz bandwidth 





R2


(390)





R1


(2K)





R3


(27)





R4


(5.1)





e





b





c





C1








C2








R6


(294)





Q1


2N 5109





3 dB Pad





R8


(294)





R7


(17.4)





T1





C3








C4








C5








R5


(27)





+15 VDC





RF Out





RF 


In





Note: C1, C2, C3, C4, C5 are 0.1 fd





T1 is 10 bifilar turns of #28 wire on Amidon FT37-43 core (4:1 Transformer)





Heat Sink required on Q1 (2N5109), current 50 ma











10.7 MHz 2N-5109 Amp and 3 dB Pad





+15 VDC





RFOut





RFIn





c





b





e





Q1





R5





C4





C5





C3





C2





C1





R6





R8





R7





R3





R4





R2





R1





T1


























































































































































































































































































































































































































Log Detector, Low Pass Filter, DC Amp with Variable Gain and Offset





Note  C2, C3, C4, C5, C8, C9, C10, C11, C13, C14 are 0.1fd














TP Array Antennas, Filters, Preamps, Power Splitters





50


ohm





50


ohm





e





u





K3





K2





K1





e





e





u





u





Noise


Generator





LCP





RCP





Programmable


Attenuator





RF In





Spectrograph





K1 is LCP select relay (energize to select LCP)


K2 is RCP select relay (energize to select RCP)


K3 is CAL/OPR relay  (energize for CAL)





During CAL both K1 and K2 are un-energized, terminating both antenna ports in 50 ohms























Spec 





Noise 





LCP 





RCP 





Mode Switch





White





Gnd





RCP/LCP





Yellow





Cal





Red





RCP





CAL





Blue





LCP





Green





64 dB





Yellow





PIC and DRIVERS





32 dB





Blue





16 dB





Orange





RS-232





8 dB





Brown





4 dB





2 dB





Purple





Power





1 dB





Red





Gnd





Wh&Blk





KAY 4450 Prog Atten





16





8





4





64





32





1





2





IC-1





IC-2





Note:	C1, C2, C3, C4 are 1 fd


	IC1 = MAX-232,  IC2 = 16F-873,  Y1 = 18.432 MHz


	Ground RESET line to reset PIC





IC-2





IC-1





2





3





4





5





6





7





8





1





15





14





13





12





11





10





9





16





1





2





3





4





5





6





7





8





9





10





11





12





13





14





28





27





26





25





24





23





22





21





20





19





18





17





16





15





K3





64dB





32dB





16dB





8dB





4dB





2dB





1





2





3





4





5





6





7





8





9





1dB





DB-9





+





+





+





+





C2








C1








C3








C4








Y1








C8


18 pf








R3








R1 (1K)








CAL








R/L








R2








C5


100fd








+





C6


0.1








+5 VDC








K2








K1








Reset








C7


18 pf








C9


0.1








+5 VDC








K3





64dB





32dB





16dB





8dB





4dB





2dB





+5 VDC





22





24





26





1





R1





C2





C3





C1





C4





C6





C5





C9





C8





Y1





C7





1dB





Pin 3 DB9





Pin 2 DB9





10





Reset





R/L Button





Cal Button





K1





K2





8





1





6





20





28











Calibrator PIC Controller





10K





10K





10K





510





1N914





1K





2N3904





+5 vdc





+5 vdc





2N3906





Omron SPDT Relay Coil


125 ohms





PIC





10K





10K





10K





510





1N914





2.2K





2N3904





+12 vdc





+5 vdc





2N3906








Step Atten Relay Coil


380 ohms





PIC





+12 vdc











+15





+5





+12





+12











-12





+5





Gnd





Gndd





Elenco Front Panel View











Spectrograph Panel











Power Filter


Module





TB3





TB2





TB1





GND





GND





GND





GND





+15





+15





+15





+15





Gray





Gray





Gray





Red





Red





Red





M6





M5





M2





TB3





GND





GND





GND





GND





+15





+15





+15





+15





Gray





Gray





Gray





Gray





Red





Red





Red





Red





M8





M9





M11





M14





TB2





GND





Black





+5





M3





Purple





+5





+15





GND





Power


Filter


Module





+12





Orange





M4





Black





GND





Orange





+12





-12





+12





M14





Black





GND





Blue





-12





TB1





Spectrograph power distribution – barrier strip wiring





10.7 MHz IF





70 MHz IF





Line and Phasing 


Loss


L=3dB





Figure 2    FS-200 system gains and losses as installed at UFRO








40K at antenna will produce -41dBm


at input to log-detector


(step attenuator loss set to zero dB)














40K at antenna produces 400K (-98dBm)


at input to M1











DC signal level out to ADC








Net gain in dB measured


from antenna elements





67





Pad





M13





Amp G=23dB, NF= 4dB


Pad L=3dB


Module G=20dB








Step


Attenuator











40K = -108dBm


(30 kHz BW)








Log Det





LPF





DC AMP





M14











10





UFRO





47





50





30





33





Pad





M11





Amp G=23dB, NF= 4dB


Pad L=3dB


Module G=20dB








17





24





27





18





25





9





18-30 MHz


BPF


(L=1 dB)





M1





4-Way


Splitter


L=7dB





17





1





4.5





7.5





-3.5





-3





M8





(L=7 dB)





Line Loss


L=3dB





M5





(L=7 dB)





RF Amp


AMP-77


(G=16dB


NF=3.5 dB)








M2





70 MHz


XTAL BPF


L=3dB





M7





Dpxr








Pad





M6





Diplexer L=1 dB


Amp G=16dB, NF=4 dB


Pad L=6dB


Module Gain=9dB








RF Amp


AMP-77


(G=16dB


NF=3.5 dB)














RF Amp


  (G=11dB


NF=3dB)














18-30 MHz


BPF


(L=0.5 dB)











2-Way


Splitter


L=3.5dB





10.7 MHz


XTAL BPF


L=3dB





M12





10.7 MHz


XTAL BPF


L=3dB





M10





Dpxr





Pad





M9





Diplexer L=1 dB


Amp G=23dB, NF=4 dB


Pad L=6dB


Module Gain=16dB








100





100





82





82





5-60





5-60





5-60





5-60





Low Pass 





L1





L2





L3





5-60





5-60





5-60





5-60





L4





RF


IN





L5





0.1





0.1





0.1





220





220





ERA-4





RF OUT





High Pass 





L6





+12 vdc





Ferrite Bead





L1 - L6 all 6 turns #21 on T50-10 core





Due to a PC board layout error pins 2 and 3 of the LF-411s have been cross - jumpered





RF


IN











R2





R1





R3





R4





R5





R7





R15





R8





R9





R13





R12





R16





R6





R17





R14





C1





C2





C3





C4





C5





C6





C8





C9





C7





IC1





IC2





IC3





R10





C15





C14





VR2





VR1





C12





C10





C13





C11





ZD1





VR3





Signal


Out





+12 VDC





-12 VDC





GND





R11





R18





Offset





Gain





Max Level Set





R1


(51)





R2 (470)





R3 (5.1)





R4 (100K)





R5 (100K)





R6 (5.1)





R7


(100K)





R8


(56K)





R9 (50K)





R10


(50K)





R11


(1K)





R12


(5.6K)





2





3





4





7





6





6





7





4





2





3





R16 (100)





1











4





5











8





C1


(0.01)








C2








C3








C4








C5








C10








C7


(820)





C6


(820)





78L05





78L09





79L09





IC1





IC2





IC3





C8








C9








C13








+9 V








-9 V








VR1








VR2








VR3








+12 VDC








-12 VDC








R17


(1K)





ZD1


1N5231





R13


(330K)





R14


(100K)





OFFSET





+





+





-





-





-





+





+





-





DC GAIN





LEVEL


SET





C11








C12


(10








C14








Signal


Out





RF


In





AD8307





LF411





LF411





R18


(1K)





Log Detector





Low Pass Filter





DC Amplifier





C15


(10








+








+








RF Associates LLC  


Sept 10, 2003





Log Detector, Low Pass Filter, DC Amp with Variable Gain and Offset





Note  C2, C3, C4, C5, C8, C9, C10, C11, C13, C14 are 0.1fd





-9 V








Min Gain





Max Gain








PAGE  
2

_1131561408.xls
Chart1

		-75		-75

		-70		-70

		-65		-65

		-60		-60

		-55		-55

		-50		-50

		-45		-45

		-40		-40

		-35		-35

		-30		-30

		-25		-25



10.7 MHz RF Input (dBm)

DC Output (Volts)

Log Detector Response

0.57

0.36

0.72

0.51

1.11

0.84

1.77

1.38

2.53

1.97

3.28

2.55

4

3.16

4.71

3.77

4.95

4.32

5

4.8

5

4.95



Sheet1

		-75		0.57		0.36

		-70		0.72		0.51

		-65		1.11		0.84

		-60		1.77		1.38

		-55		2.53		1.97

		-50		3.28		2.55

		-45		4		3.16

		-40		4.71		3.77

		-35		4.95		4.32

		-30		5		4.8

		-25		5		4.95





Sheet1

		



10.7 MHz RF Input (dBm)

DC Output (Volts)

Log Detector Response



Sheet2

		



10.7 MHz RF Input (dBm)

DC Output (Volts)

Log Detector Response



Sheet3

		





		






